## Sheet1
Program: [ 5.1.0.0 - RP2 - X-ray sources driven by rep-rate ultrashort laser pulses ]
Subject: [ large size spare Spherical mirror for Long focal length setup in E3 ]. Verification Control Document (VCD)
Specification: [ TC ID # 00260975/C - RSD_High Quality, large size spare Spherical mirror for Long focal length setup in E3_SP20_002 ]
RSS TC ID # 020260/A.003;High Quality, large size spare Spherical mirror for Long focal length setup in E3_SP20_002
Requirement TC ID | Rev. | Requirement text | Verification Method | VM clarification (see also QRs, REQ-028651/A) | Close-out | Verification Record Document (ID # QR or protocol, declaration, etc., page # ) | Comments
Yes / No
2. General requirement
REQ-028617 | A | The Supplier shall deliver 1 pc of the Spherical mirror in accordance with requirements described in the subchapter 3. | I – inspection
3. Functional, Performance and Design requirements
3.1. Geometry and Substrate
REQ-028618 | A | The substrate material of the SP mirror shall be manufactured from mirror grade ULE, Zerodur, fused silica, high-power laser compatible or equivalent quality material. | R - review | Substrate material report or certificate
REQ-028619 | A | The SP mirror (along the beam) shall be rectangular in shape with physical dimension of (310+0-1) x (270+0-1) mm2. NOTE: The dimensions shall correspond to the requirements given in the reference drawing RD-01 (see chapter 1.4) which is an inseparable part of the SP mirror requirements. | R - review, T - test | Dimensional report
REQ-028620 | A | The clear aperture of the SP mirror (along the beam) shall be squared (center of the clear aperture is the mirror center). The clear aperture shall be squared with dimensions at least 250x250 mm2 and with rounded corners. | R - review, T - test
REQ-028621 | A | The center thickness of the SP mirror shall be (60±1) mm. | R - review, T - test | Dimensional report
REQ-028622 | A | The radius of curvature (center of the SP mirror to the focal point) shall be (10000±20) mm. | R - review, T - test
3.2. SP mirror coating
REQ-028625 | A | The front (i.e. reflective) surface of the SP mirror shall have a protected Au coating. | R - review | Coating material report
REQ-028626 | A | The damage threshold of the SP mirror coating shall be at least 100 mJ/cm2, 1000-on-1 @ 10 Hz @ 30 fs (central wavelength 810 nm). | R - review, A - analysis
REQ-029485 | A | Coating adhesion of the SP mirror shall be in accordance with MIL-C-48497A. | R - review
3.3. SP mirror surface
REQ-028627 | A | The reflectivity of the SP mirror shall be not less than 95 % over the bandwidth of Δλ ranging with 766 nm ≤ Δλ ≤ 854 nm. NOTE: The Supplier shall use its best efforts to provide the reflectance homogeneity over the clear aperture of the SP mirror. | R - review, T - test | Reflectivity measurement report
REQ-028628 | A | The surface accuracy of the SP mirror shall be maximum λ/10 peak-to-valley over at least 240 mm x 240 mm under test wavelength of (633±2) nm. | R - review, T - test | Interferometric report
REQ-028629 | A | The RMS reflected wavefront of the SP mirror shall be lower than λ/20 (Power term removed). | R - review, T - test | Interferometric report
REQ-028630 | A | The RMS surface slope error of the SP mirror (after coating) shall not exceed 2.5 µrad (i.e. 30nm/1cm) with 1 to 10 mm as spatial periods applying to the slope metrology. | R - review, T - test | Interferometric report
REQ-028631 | A | The surfaces quality of the SP mirror shall not exceed 60/40 scratch/dig according to the MIL-PRF-13830B standard. NOTE: The Supplier shall use its best efforts to reach better surface quality, i.e. 40/20 scratch/dig. | R - review, T - test | Surfaces quality report
REQ-028632 | A | The Group Delay Dispersion (¦GDD¦) of the SP mirror shall not exceed 30 fs2 for a bandwidth of Δλ ranging with 766 nm ≤ Δλ ≤ 854 nm for p-polarization. | R - review, A - analysis | ¦GDD¦ estimation report
REQ-028633 | A | The microroughness of the SP mirror shall not exceed 2 nm (after coating). | R - review, T - test | Interferometric report
REQ-028634 | A | The 2 pcs 2-inch diameter substrates manufactured from the exact same material and polished with the exact same surface finish as the SP mirror shall be coated together with the SP mirror and provided to the CA as a LDT witness sample for laser damage testing at the ELI Beamlines or any other certified/qualified institution and to measure the spectrum of the reflectivity and the GDD. | R – review, I - inspection
3.4. Other SP requirements
REQ-028636 | A | The SP mirror shall be vacuum compatible and shall be able to be totally operated under vacuum level up to 10-6 mbar. In this working environment, the SP mirror shall comply with all specifications given in the subchapter 3.1 and shall not be outgassing. | R - review, A - analysis
REQ-028637 | A | The chamfers of the SP mirror substrate shall comply with the reference drawing RD-01 (see chapter 1.4). | R - review, T - test | Dimensional report
REQ-028638 | A | The lateral side of the SP mirror shall have the mounting interface in the form of a groove along 2 vertical sides as defined in in the reference drawing RD-01 (see chapter 1.4) which is inseparable part of the SP mirror requirements. The shape of base of the grooves shall be flat. | R – review, I - inspection
REQ-028639 | A | The SP mirror shall have the following information engraved into the top of the substrate barrel: □ Part number according to Figure 1 below: ▼ ELI EX L3 SPM 0810 PQQAA LLLxWWWxTTT XXXXX ▼ □ Mark (triangle) pointing to the HR coated front side S1. NOTE1: The following example demonstrates the correct engraved information: ▼ ELI E3 L3 SPM 0810 PQQ00 310x270x060 00001 ▼ NOTE2: This engraved information (see NOTE1 above) denotes as follows (see Figure 1 below): • ELI Beamlines, • Experimental hall 3, • L3 laser Spherical mirror @ 810 nm, • p-polarization, 0A quality, • Substrate size, • Serial number, • Triangles pointing to HR coated side S1. | I – inspection
4. Environmental requirements
REQ-028645 | A | The Supplier and the CA shall agree on the cleaning method to clean the SP mirror without decreasing the mirror’s properties and to avoid contamination of clean space. NOTE: The cleaning methods may use high gas flow (dry and clean air or N2) and specialized chemical cleaning liquids (alcohol, Isopropyl alcohol, demineralized water). | R - review
5. Delivery requirements
REQ-028646 | A | The transportation to the final destination of the SP mirror shall be conducted by the Supplier. NOTE: The Supplier is fully responsible for the delivery of undamaged SP mirror. | I – inspection
REQ-028647 | A | The SP mirror shall be delivered in a protective container – preferably a Polyethylene Terephthalate Glycol (PETG) container – that shall not contact the coated CAP of the SP mirror and prevents its damage, degradation and contamination. The container shall then be sealed with ultra-low outgassing clean room tape and wrapped into 2 aluminized clean room 100 compatible polythene bags. Packaging shall be performed in an ISO class 6 or better clean room environment 6 according to ČSN EN ISO 14644 (or equivalent, e.g. EN ISO 14644). | I - inspection, R - review
REQ-028648 | A | The Supplier shall provide an appropriate shipping crate with shock and tit watch and packaging material compliant for overseas shipping. | I – inspection
REQ-028649 | A | The serial number shall be clearly marked on the outer layer of the bubble wrap. The aluminized bag shall have the label: “To be opened only in a clean room class ISO 6 or better and by trained personnel as well as "FRAGILE – GLASS". NOTE: The same labels should be applied on the outside of the shipping crate/box. | I – inspection
6. Safety Requirements
REQ-028650 | A | The Supplier shall supply a Declaration of Conformity or any other equivalent document legally recognized and accepted in the Czech Republic for each product type if the appropriate legislation determines the Supplier's obligation to have a Declaration of Conformity (or the equivalent document) for the purposes of a Device sale in the Czech Republic to fulfil the requirements of 2001/95/EC directive or applicable Czech law. | I – inspection
7. Quality control
7.1. Quality Reports (QRs)
REQ-028651 | A | The Supplier shall perform a factory verification of the SP mirror and provide corresponding specific quality reports (I - VII) proving that the requirements have been met: I. Interferometric report before and after coating of the SP mirror surface with a map showing a departure from the ideal reflected wavefront in units of testing wavelength λ and/or nm as well as wavefront RMS, gradient and mid spatial frequencies (see REQ-028628/A, REQ-028629/A, REQ-028630/A and REQ-028633/A). II. Reflectivity measurement report of the central area over the specified bandwidth Δλ (see REQ-028627/A) or of the witness sample at ambient humidity as well as at < 2% RH. III. GDD measurement or estimation report. NOTE: The report may be carried out by the Supplier via measurements performed or calculation or simulations including the description of the estimation method (see REQ-028658/A). IV. Surface quality report with scratch and dig map showing the locations of all defects (see REQ-028631/A). V. Substrate material report or certificate showing what is the material of the substrate (see REQ-028618/A). VI. Coating material report or certificate showing what is the material of the coating (see REQ-028625/A). VII. Dimensional report providing information about a measured physical dimension of the manufactured product (see REQ-028619/A, REQ-028621/A and REQ-028637/A). NOTE 1: The results of the factory verification of the SP mirror shall be provided to the CA in the corresponding specific QRs before mirror delivery. NOTE 2: The results shall be in print as well as in digital format. | R – review, I - inspection
7.2. Documentation and data control
REQ-028652 | A | The Supplier shall provide Instructions for use (Product User Manual) as part of the delivered Product. The manual shall include the instructions and descriptions regarding the following:transport, handling, storage and cleaning; safe operation and maintenance procedures. | R – review, I - inspection
REQ-028653 | A | The Supplier shall supply the following relevant manufacturing documents:all approved by the CA manufacturing drawings, 3D models and design supporting documentation; all approved by the CA “requests for deviation/wavier from requirements described herein”; NOTE: The manufacturing documents may be supplied in electronic form in relevant data formats described in the REQ-028656/A. | I – inspection
REQ-028654 | A | The manufacturing documents shall contain strictly the units which are used to define the requirements in chapter 3. | R - review
REQ-028655 | A | The manufacturing documents shall include the accuracy of the manufacturing process. This accuracy shall be also included in the corresponding QRs (see REQ-028651/A). NOTE: The Supplier will specify what can be the maximal difference between the specified parameters (in chapter 3) and the parameters of the final product. | R - review
7.4. Verification requirements for the Supplier
7.4.2. Verification Control Document (VCD)
REQ-028659 | A | The Supplier shall provide a Verification Control Document (VCD) for the reviews as agreed with the CA. NOTE 1: Guidelines for VCD preparation will be provided by the CA (see chapter 1.4; RD-02). NOTE 2: The form of VCD will be agreed between the CA and the Supplier based on the best commercial praxis used by the Supplier. NOTE 3: The form of VCD will be agreed before completion of the Qualification Design phase (see chapter 7.5.1). NOTE 4: The VCD specifies HOW and WHEN each requirement is planned to be verified by the Supplier, when it was actually verified. | R - review
7.5. Phasing of the delivery
7.5.1. Qualification of Design
REQ-028660 | A | Before completion of the Qualification Design phase the Supplier shall provide following information that shall be agreed by the CA:structure and content of the Quality Reports (QRs, see REQ-028651/A);structure and content of the VCD ready to be implemented (see REQ-028659/A). | R - review
REQ-028661 | A | Before completion of the Qualification Design phase the Supplier and the CA shall agree on:final manufacturing drawings provided by the Supplier (see REQ-028653/A);detailed procedures related to the testing during Manufacturing phase (see chapter 7.5.2 Manufacturing);common nonconformity control system (see REQ-028657/A). | R - review
7.5.2. Manufacturing
REQ-028662 | A | The results of the Manufacturing phase of verification shall be recorded by the Supplier in the appropriate QRs (or in other factory/quality reports, if not specified in chapter 7.1) and overall results (including review of documentation/reports and inspection of product) shall be recorded in the VCD (see chapters 7.4.2). NOTE: The results of all the tests shall be given strictly in units which are used to define the requirements in chapter 3. | R - review
REQ-028663 | A | The final issue of the VCD shall be submitted to the CA after the approval of the last report and before starting the acceptance (see chapter 7.5.3). | R - review