ZD Annex No. 1 - Technical Specification - Supply of centrifugal partitioning chromatograph with accessories

General information on design and engineering requirements

The following media are available for the technology:

- cooling water T=15o C

- waste for max. T=100o C

- electrical energy approx. 50 kW

- air handling

The media are introduced into the production facility in the form of connection points.

Environment:

According to the preliminary protocol for the determination of external influences according to CSN 33 2000 5-51 for the production plant, the following basic external influences are determined:

AA5 - temperature +5 to +40°C, relative humidity 15-100%

AF2 - occurrence of organic solvents

BA4 - Capabilities of Persons (Instructed Persons)

Other environmental influences are negligible. The only exception is the site at the distillation plant where it is determined:

AA6 - elevated temperature limit

The size of the space for the placement of the technology:

Maximum 3 890 mm x 3 400 mm

The contracting authority requires the supply and installation of complete centrifugal partition chromatography technology consisting of the following equipment:

1. CPC including all accessories necessary for its operation (injection port, high pressure pump, PDA detector, multiway valve for fraction collection, control software)

2. Rotary vacuum evaporator, 2 pcs

3. Solvent system preparation module (mobile LP and stationary UP phase) for CPC:

3.1. Mixing tank

3.2. Settling tank (settler)

3.3 Upper (UP) and Lower (LP) phase storage tanks

3.4. Filter for water separation

4. Density measurement

5. Gas chromatograph

All components of the process equipment must be equipped with the necessary measuring and control elements including their necessary accessories. The individual devices form separate technological units (units) and can be installed, commissioned, handed over and taken over separately. A diagram of each device shall be included in the tender.

Specification of individual devices (technological nodes)

1. Centrifugation partition chromatograph (CPC)

The CPC plant will be used for the preparation of pure fractions from crude extracts of autotrophic micro-organisms (in particular the green alga Haematococcus pluvialis). The key parameter for the planned application is a high separation efficiency in the shortest possible time, with a minimum of 100 grams of raw material (extract) per injection. For this reason, we require a hydrostatic equilibrium system with a constant gravitational field, which will be generated by rotation of the separation column (rotor) in one axis. The equipment must be capable of operating at an operating pressure in the range of 2-7 MPa (300-1,000 psi) and a rotor speed in the range of 300-700 rpm. The rotor volume shall be a minimum of 10 litres. The CPC shall be capable of operating in two operating modes, ascending and descending. In ascending mode, the lighter/upper phase (UP) shall be the mobile phase and in descending mode, the heavier/lower phase (LP) shall be the mobile phase. The equipment shall be capable of changing the direction of flow of the mobile phase during the process and thus operate in dual mode, i.e. in ascending mode combined with descending mode. The CPC rotor and other parts (rotor chassis and other structural elements) shall be made of AISI 316 stainless steel with an internal surface of Ra < 0,8 μm, PTFE or PTFE coated silicone seals shall be used. Material in contact with organic solvents AISI 316, with internal surface Ra < 0,8 μm. The instrument shall have a control panel allowing adjustment and measurement of rotation speed, temperature and air flow through the stainless steel exhaust pipe outside the ATEX I area. The CPC shall also be equipped with mobile/stationary phase leak detection, a volatile solvent vapour sensor and an excessive rotor vibration sensor. The rotor access door of the CPC shall have a safety locking mechanism to prevent it from opening if the rotor is moving. The CPC shall be designed to operate on a standard EU 220 to 240 V, 50 Hz mains supply. The CPC shall comply with the ATEX I Directive. The CPC separator itself shall also contain all the additional equipment necessary for its operation, namely the injection port with all necessary accessories (valves, spare sets, additional valves and tools), a high-pressure gradient pump with a flow rate of up to 1000 ml/min with the possibility of pumping the upper or lower phase and also of injecting the sample, at least a 10-position computer-controlled switching valve for fraction collection, one UV-Vis detector with photodiode array (PDA), an A/D converter for an additional independent detector (e.g. ELSD), personal computer and control software. The software shall be capable of controlling all accessories, i.e. the high pressure gradient pump (flow rate setting in the range of 5-1000 ml/min. and gradient creation), sample injection, fraction collection diverter valve and data acquisition from the photodiode UV-Vis detector. The tender shall also include all hoses, connections and fittings, spare parts, installation and on-site training.

2. Rotary vacuum evaporator (RVO) - 2 pieces

For the recycling of the mixture of organic solvents from the isolated fractions (Acetonitrile, Heptane, Hexane, Ethanol, Ethyl acetate, Acetone) after separation by CPC, a bid for a fully automated large capacity Rotary Vacuum Evaporator (RVO) will be accepted. The RVO must allow for fully automatic distillation with an evaporating flask of at least 20 litres capacity. The RVO shall include in the basic assembly a control unit allowing control of vacuum, temperature, speed, water level in the heating bath and automatic shutdown if the required pressure is exceeded or the glass is damaged. The RVO shall be supplied with a minimum of 2 evaporation flasks of 20 litres capacity, 2 evaporation flasks of 10 litres capacity, 1 condenser, 1 vacuum pump of adequate capacity and 1 chiller. The speed of the evaporating flask must be adjustable between 10-150 rpm, the temperature range of the bath when using water 20-100 °C, the lifting of the flask is possible manually or by means of a motor, the heating capacity of the bath is at least 4 kW, the power supply is 400 V.

3. Solvent system preparation module (UP/LP)

In order to recycle solvents and reuse them for CPC separation, the composition of the UP and LP phases must be maintained in the required ratio. Off-line measurement of the density of the separated phases and gas chromatography will be used to accurately measure the composition of the solvent system (UP/LP). The results of the rapid analysis of the UP and LP composition after distillation will be used to accurately adjust the composition of the solvent mixture manually.

The mobile phase preparation module consists of a mixing tank (mixer), a settling tank (settler), 2 tanks of pre-distilled and treated mixture of the upper (UP) and lower (LP) phase, a centrifugal pump, control valves and the necessary interconnecting piping (see Fig. 1. below).

The mixtures of solvent systems (UP and LP) used for CPC separation are fed into a stirred tank to ensure that they remain homogeneous prior to distillation. From the stirred tank, the mixture is continuously fed to a fully automated large capacity rotary vacuum evaporator where it is pre-distilled as a mixture. The recycled solvent system mixture is pumped from the distillation apparatus into a settling tank (settler), where it is separated again into upper (UP) and lower (LP) phases. Based on compliance or deviation of the desired phase composition from the desired value, as determined by rapid density analysis and gas chromatography, the solvent system is manually replenished with the required amount of solvents. Once the correct phase ratio has been set, the solvent system (UP/LP) is fed via a centrifugal pump into the UP or LP phase reservoirs. The UP and LP phases thus prepared will be stored in storage tanks from which they will be pumped to the CPC and used for subsequent separation.

3.1. Mixing tank (mixer)

Equipment for storage of the used solvent system (UP or LP) required for separation by CPC with mixing capability. A freestanding, single-walled, cylindrical, non-insulated, non-pressurised, stainless steel vessel with a working volume of at least 100 litres, with a hinged top lid and a sight glass of at least 100 mm diameter, fitted with a speed-controlled agitator with the necessary controls and metering and regulation elements. The hinged top lid shall be fitted with a seal to prevent the escape of solvent vapours. All structural elements, measurement and control, valves, inlet and outlet pipes shall be mounted on the same support frame as the mixer. The support frame is fitted with locking castors. The castors shall have adequate load capacity to withstand the load of a full mixer and shall be capable of being wheeled to move the mixer on a level surface or to turn it on the spot. At the same time, the wheels must be arranged in such a way as to ensure the stability of the mixer against tipping, both when travelling and when standing still. The maximum height of the whole device, including the open lid (counting the lid at its highest point, i.e. when the assembly is at its highest), must not exceed 2000 mm and the device must not exceed a base of 500 x 500 mm.

All parts of the mixer in contact with organic solvents must be inert (AISI 316 stainless steel) with a smooth surface (at least Ra 0.8 μm) with no dead spaces for the possibility of depositing impurities. Other structural parts in AISI 304 (Ra < 1.6 um), PTFE or PTFE coated silicone seals.

Mixing by means of a propeller or paddle mixer connected by a shaft to a motor located in the mixer lid parallel to the axis of the container. Drive the mixer with speed control in the range 0 - 750 rpm. Stops shall be fitted on the inner wall of the mixer to prevent the liquid from rotating during mixing and the formation of a central vortex. The mixing medium shall be close to water in viscosity and shall be free of solid abrasive particles (sand, alumina, silica, limestone, etc.).

The bottom of the vessel shall also provide for the discharge of the vessel from the lowest point of the vessel. The vessel shall be fitted with a level limit sensor to prevent overfilling of the vessel. The agitator drive motor shall be ATEX category II explosion-proof.

3.2. Settling tank (settler)

The sedimentation tank shall be made of chemically resistant borosilicate glass with a minimum volume of 100 litres and a DN25 stopcock in the bottom of the tank. The tank shall be covered with a flat universal lid DN 400 with three openings to allow filling of phases, mixing and sampling for measurement of phase composition. Sealing of the lid with double-sided profiled PTFE DN 400. The glass settling tank shall be fixed in a stainless steel support frame. The support frame shall be fitted with locking castors. The castors shall have adequate load capacity to withstand the load of the glass tank including the solvent, and shall be capable of being wheeled to move it on a flat surface or to turn it on the spot. At the same time, the wheels must be so arranged as to ensure the stability of the tank against overturning, both when travelling and when standing still. Requirement for mixing by means of a circulation loop.

3.3 Upper (UP) and Lower (LP) phase storage tanks

All necessary containers for the storage of UP and LP shall be made of chemically resistant material.

3.4. Filter

The client requires a filter for separation of water from solvent mixtures based on silica gel or molecular sieve. The cartridge shall be replaceable and recyclable.

4. Density measurement

Laboratory digital hydrometer with accuracy to at least 3 decimal places. The maximum sample volume for a single measurement must not exceed 2 ml.

5. Gas chromatograph

Gas chromatograph equipped with Split/Splitless injection with electronic pressure control of service gases in the range of at least 0.5 - 900 kPa, flow rate of at least 1 - 1200 ml/min. The oven of the gas chromatograph shall be capable of a minimum of 20 temperature ramps with oven temperature settings of at least 50-400°C with a maximum heating rate of up to 50°C/min. The gas chromatograph shall be equipped with a flame ionisation detector (FID) at the base with electronic control of the pressure of the service gases, data acquisition rate of at least 250 Hz. The delivery must include a PC (PC with LCD monitor, keyboard and mouse) and software for controlling the gas chromatograph, setting up GC methods and evaluating results, Microsoft Windows operating system. The instrument shall allow for future expansion of the gas chromatograph with an autosampler. The tender price must include all costs, i.e. incidental costs such as transport, removal and disposal of packaging, assembly at the installation site, training of the customer's staff.

Fig. 1. Schematic of the solvent system preparation module (UP/LP).