## tender
SEM-EDS-EBSD |  |  | High-resolution scanning electron microscope with analysis of element and phase composition and crystallographic orientation (translation of the tender title from Czech) | required |  | option |  | max |  | POINTS | BID |  | Notes
|  |  |  | yes / value | unit | yes | scoring | 11790 |  | 0 | 0%
|  |  | All parameters must be guaranteed and documented either in the instrument specification data sheet or by a specific document. They must conform to current configuration on the market. |  |  |  |  | 7640 | SEM | 0 | 0%
[8 empty] | 2850 | ANAL | 0 | 0%
[8 empty] | 1300 | SUPP | 0 | 0%
[11 empty] | number or | unit
1-yes / 0-no
MICROSCOPE (SEM) | [9 empty] | 0
optics
1 |  | electron source | FEG (Schottky type) with minimum lifetime warranty of 1 year and first replacement included | 1 | year
2 |  |  | maximum HT | 30 | kV
3 |  |  | minimum landing HT | 50 | V | yes | 5 p / 1 V | 200 | 5 |  |  | V |  | x
4 |  | beam | minimum beam current [pA] | 3 | pA | yes | 20 p / 1 pA | 50 | 20 |  |  | pA |  | x
5 |  |  | maximum beam current at 30 kV [nA] under imaging conditions | 100 | nA | yes | 5 p /10 nA | 200 | 0.5 |  |  | nA |  | x
6 |  |  | beam current stability / 10 h [%] | 1 | % | yes | 50 p / 0,1% | 300 | 500 |  |  | % |  | x
7 |  |  | beam deceleration |  |  | yes | 200 p | 200 | 200
8 |  |  | - possibility of positive bias at least 600 V |  |  | yes | 100 p | 100 | 100
9 |  |  | integrated beam current measurement | yes
10 |  |  | - Farraday cup |  |  | yes | 200 p | 200 | 200
11 |  |  | - other probe current meter |  |  | yes | 20 p | 20 | 20
12 |  | column | column specifics (extra accelaration) |  |  | yes | 300 p | 300 | 300
13 |  | objective lens | electrostatic and field-free magnetic lens | yes
14 |  |  | immersion magnetic lens |  |  | yes | 200 p | 200 | 200
15 |  | magnification | magnification range from 15 x to 1 000 000 x | yes
low vaccum
16 |  |  | low vaccum adjustable in the range 10 - 100 Pa (the sum of Pa below and above the required range is added to column L, the actual range offered is added to column N) | 10 - 100 | Pa | yes | 0,5 p / ±1 Pa | 200 | 0.5 |  |  | Pa
17 |  |  | low vaccum achieved by using nitrogen or water vapor (what you use => nitrogen = 0; water vapor = 1) |  |  | yes | 50 p / water vapor | 50 | 50
resolution
18 |  | HiVac mode | 15 kV (SE) | 0.9 | nm | yes | 50 p / 0,1 nm | 200 | 500 |  |  | nm |  | x
19 |  |  | 1 kV (SE) | 1.3 | nm | yes | 50 p / 0,1 nm | 300 | 500 |  |  | nm |  | x
20 |  | LowVac mode | LV at 30 kV [nm/Pa] - @200 Pa | 2 | nm | yes | 20 p / 0,1 nm | 300 | 200 |  |  | nm |  | x
21 |  | STEM | STEM [nm/30 kV] | 1 | nm | yes | 20 p / 0,1 nm | 150 | 200 |  |  | nm |  | x
detectors
22 |  | in-beam | in-beam SE and BSE detectors | yes
23 |  |  | - energy filtered (additional) |  |  | yes | 100 p | 100 | 100
24 |  |  | - number of segments |  | segment | yes | 100 p / segment | 400 | 100 |  |  | segment(-s)
25 |  |  | - additional SE in-beam detector |  |  | yes | 500 p | 500 | 500
26 |  | other | SE detector (Everthart -Thorney) | yes
27 |  |  | directional BSE (segmented) - below OL with automatic insertion |  |  | yes | 500 p | 500 | 500
28 |  |  | - number of segments |  | segment | yes | 50 p/ segment | 350 | 50 |  |  | segment(-s)
29 |  | LowVac mode | low-vacuum SE detector | yes
30 |  |  | low-vacuum BSE detector | yes
31 |  | STEM | retractable STEM detector with BF, DF and HADF capability | yes
32 |  |  | - number of segments |  | segment | yes | 50 p / segment | 550 | 50 |  |  | segment(-s)
chamber
32 |  |  | Minimum number of ports | 10 | port | yes | 5 p / port | 50 | 5 |  |  | port(-s)
33 |  |  | intergrated decontaminator / plasma cleaner | yes
34 |  |  | compatibility with FemtoTools - 3 cabels | yes
35 |  |  | airlock / loadlock |  |  | yes | 50 p | 50 | 50
36 |  | IR camera | for navigation and safety | yes
37 |  | navigation camera | color optical camera with possibility to save image | yes
38 |  |  | image anotation |  |  | yes | 50 p | 50 | 50
39 |  | collision protection | automatic collision protection with touch alarm | yes
40 |  |  | - using sample holder and sample dimensions |  |  | yes | 100 p | 100 | 100
41 |  |  | - including SEM detectors |  |  | yes | 100 p | 100 | 100
42 |  |  | - including analytics (EDS + EBSD), third party |  |  | yes | 50 p | 50 | 50
motorized stage
43 |  |  | either eucetric or compucentric with store & recall, click-to-center, drag-to-zoom, and image align to horizontal/vertical functions | yes
44 |  |  | eucentric |  |  | yes | 200 p | 200 | 200
45 |  |  | compucentric |  |  | yes | 200 p | 200 | 200
46 |  | stage movement | XY range minimum 50 x 50 mm | 50 x 50 | mm x mm | yes | 200 p / min. 100 x 100 mm | 200 | 200 |  |  |  |  | x
47 |  |  | Z range minimum 40 mm | 40 | mm | yes | 200 p / min. 60 mm | 200 | 200 |  |  |  |  | x
48 |  |  | full 360 deg rotation with tilt minimum -4 to +70 deg | -4 to +70 | deg | yes | 200 p / +90 deg | 200 | 200 |  |  |  |  | x
scanning
49 |  | field of view (FOV) | FOV at WD for EDS minimum of 3 mm | 3 | mm | yes | 5 p / 1 mm | 70 | 5 |  |  | mm
50 |  | resolution | minimum resolution 6144 x 4096 pixels and 360 deg rotation | 6144 x 4096 | pixel
51 |  | correction | shift, scan, tilt , and rotation | yes
software
52 |  | access | muli-user acces / accounts and user / admin permissions | yes
53 |  |  | observation recepies for various samples |  |  | yes | 100 p | 100 | 100
54 |  | auto | focus, BC, stigmation, lens alignment | yes
55 |  | drift | drift correction | yes
56 |  | imaging | live and scanning strategies to reduce exposure (e.g. line, frame, interlaced, serpentine accumulation) | yes
57 |  |  | button "undo" |  |  | yes | 100 p | 100 | 100
58 |  | number of signals viewed simultaneously with possibility of mixing signals |  | 4 | signal | yes | 50 p / signal | 200 | 50 |  |  | signal(-s)
59 |  | presets | storing column and scanning setting for further use (number of presets) | 4 | preset | yes | 50 p / preset | 200 | 50 |  |  | preset(-s)
60 |  | save | image format as tiff / bmp / jpg (which are required) with databar or image only (additional formats shall be listed in column N) | tiff / bmp / jpg | format | yes | 20 p / type | 40 | 20 |  |  | format(-s)
61 |  |  | dynamic range (8 / 16 / 24 bit) | 8 /16 bit |  | yes | 20 p / 24 bit | 20 | 20
62 |  |  | movies format as avi (which is required, additional formats shall be listed in column N) | avi | format | yes | 20 p / type | 20 | 20 |  |  | format(-s)
63 |  |  | with large, well legible scalebar only |  |  | yes | 50 p | 50 | 50
64 |  | reports | creating reports from measurement |  |  | yes | 20 p | 20 | 20
65 |  |  | template format adjustsment |  |  | yes | 10 p | 10 | 10
66 |  | large areas | automatic acquisition of image mosaics including auto focuss and CB adjustmen at every position | yes
67 |  | scripting | automation by scripting in python | yes
68 |  | licences | number of off-microscope licenses (e.g. for image export and reporting) |  | license | yes | 20 p/ license | 40 | 20 |  |  | license(-s)
accessories
69 |  | PC | high-end PC with large storage to handle acquisition, storage and processing and monitor(s) at least 24 inch with minimum resolution of 1920 x 1200 | yes
70 |  | sample move | by keyboard and mouse plus trackball or joystick (to be filled in column N) | joystick / trackball
71 |  | control pannel | including adjustment knobs or buttons for magnification, constrast, brightness, beam shift, stigmation, multifunction knobs | yes
72 |  | remote access | remote access for users and for diagnostics | yes
73 |  | hardware | antivibration system | yes
74 |  |  | oil-free pumping | yes
75 |  |  | oil-free air compressor | yes
76 |  |  | water chiller if needed | yes
ANALYTICS | [9 empty] | 0
EDS-EBSD system
77 |  | compatibility | EDS-EBSD system compatible with SEM allowing for beam and stage control | yes
78 |  |  | separate and simlutaneous acquisition of EDS and EBSD | yes
79 |  |  | possiblity to use in LowVac mode |  |  | yes | 100 p | 100 | 100
80 |  |  | high-end PC with large storage to handle acquisition, storage and processing | yes
| EDS-EBSD | software
81 |  |  | all required analytical features for acquisition and data processing are integrated in one software |  |  | yes | 100 p | 100 | 100
EDS
| EDS | hardware
82 |  | detector | SDD with Peltier cooling and motorized insertion controlled from SW | yes
83 |  |  | protective shutter | yes
84 |  |  | chip size at least 100 mm2 | 100 |  | yes | 50 p / 10 mm2 | 350 | 5 |  |  | mm2
85 |  |  | detection and quantification from Be to Cf with minimum energy detection at 73 eV (Al L) | yes
86 |  | resolution | energy resolution at MnK 127 eV or better | yes | [9 empty] | x
87 |  |  | resolution MnK at 200k cps (output), specify input cps and dead time provide a proof or guarantee | 140 |  | yes | 20 p / 1 eV | 300 | 20 |  |  | eV |  | x
88 |  |  | energy resolution at CK 50 eV or better | yes | [9 empty] | x
89 |  |  | resolution CK at 50k cps (output), specify input cps and dead time provide a proof or guarantee | 50 |  | yes | 20 p / 1 eV | 100 | 20 |  |  | eV |  | x
| EDS | SW - data acquisition
90 |  | acquisition of two imaging signals from the microscope |  |  |  | yes | 100 p | 100 | 100
91 |  | acquisition | point and multipoint, line, and area acquisition (rectangular or any shape) | yes
92 |  |  | live mapping (raw), including peak identification , deconvolution, and quantification, cps mapping | yes
93 |  |  | simultaneous acquisition and display of live SEM images, live EDS spectra with automatic peak identification, and live elemental distribution EDS maps in real time even with microscope stage movement and magnification changes, without interrupting or stopping data acquisition. |  |  | yes | 100 p | 100 | 100
94 |  | automation of analysis series (point, line, area) from the previously saved positions |  | yes
95 |  |  | automatic recording of the identified elements within the analyzed fields on the sample, recording the positions of the analyzed fields and displaying them on the background of the registered images | yes
96 |  |  | automation of data acquisition and analysis from larger areas with the possibility of stitching the results into one file (image) | yes
97 |  | drift correction | live drift correction | yes
98 |  |  | predictive drift correction |  |  | yes | 50 p | 50 | 50
99 |  | standards | using standards with possibility to create own library of standards | yes
| EDS | SW - data analysis
100 |  | analysis | quantitative in atomic and weight percents with normalized and not normalized sums | yes
101 |  |  | background substraction, peak identificatin, correction for escape, fluorescence, and sum peaks, ZAF correction, peak deconvolution | yes
102 |  |  | manual background correction |  |  | yes | 50 p | 50 | 50
103 |  | mapping | quantitative and phase mapping, includimg analyses of modal composition of phases (based on phase mapping) | yes
104 |  |  | adding spectra from a selectd area from a map into one spectra (to increase SNR) | yes
105 |  |  | adding spectra of the same phase from the map into one spectra (to increase SNR) | yes
106 |  | export | export EDS spectra as image (e.g. jpg, tiff) and text (e.g. txt, csv) | yes
107 |  |  | export EDS spectra as EMSA |  |  | yes | 20 p | 20 | 20
108 |  |  | export maps as image (e.g. jpg, tiff) | yes
109 |  |  | export maps as text (matrix) |  |  | yes | 20 p | 20 | 20
110 |  | licence | number of off-microscope licenses for the EDS data processing software | 2 |  |  | 10 p / license | 100 | 10 |  |  | license(-s)
EBSD
| EBSD | hardware
111 |  | camera | CMOS with phospor screen with motorized insertion controlled from SW and possible acquisition at lower HT below 5kV and low current of 100 pA and sensitivity 700 pps/nA | yes
112 |  |  | maximum resolution of 640 x 480 pxs | 640 x 480 |  | yes | 200 p / higher resolution | 200 | 200
113 |  |  | motorized camera tilt controlled from software |  |  | yes | 100 p | 100 | 100
114 |  |  | maximum acquisition speed at least 5700 pps (give pixel resolution) | 5700 |  | yes | 100 p / 1000 pps | 100 | 0.1 |  |  | pps
115 |  |  | integrated imaging (topography, composition, orientation constrast) | yes
116 |  | collision protection | integrated collision protection | yes
117 |  |  | pre-touch sensor |  |  | yes | 100 p | 100 | 100
118 |  | transmission EBSD | Transmission Kikuchi Dffraction (TKD) together with holder | yes
119 |  |  | near-axis TKD together with holder |  |  | yes | 50 p | 50 | 50
| EBSD | SW - data acquisition
120 |  | acquisition | point and area acquisition including background and distortion correction | yes
121 |  |  | possibilty to save EBSD patterns during acquisition | yes
122 |  | structure definition | import cif-file or define own structure (unit cell, SG, element and symmetry position) | yes
123 |  |  | import from integrated database (AMCSD and ICSD required, additional ones shall be listed in column N) | AMCSD + ICSD |  | yes | 100 p | 100 | 100
124 |  | on-microscope indexing | Hough transform plus basic indexing including measure of data / inexing quality | yes
| EBSD | SW - data treatment
125 |  | structure definition | import cif-file or define own structure (unit cell, SG, element and symmetry position) | yes
126 |  |  | import from integrated database |  |  | yes | 50 p | 50 | 50
127 |  | basic indexing method | Hough transform plus basic indexing including measure of data / inexing quality | yes
128 |  | advanced indexing | template matching - dictionary indexing | yes
129 |  |  | advanced pattern matching method (using dynamically simulated patterns) | yes
130 |  |  | by projection of patterns on the sphere |  |  | yes | 500 p | 500 | 500
131 |  | phase ID and orientaion with EDS analysis support |  | yes
132 |  | display results | orientation by pole figures, inverse pole figures, Euler angles with angular precision below 0.05 deg with possibility of changing reference system of pole figures | yes
133 |  |  | particle size, Kernel average misorientation | yes
134 |  | advanced features | averaging of neighbours |  |  | yes | 50 p | 50 | 50
135 |  |  | taking into account the neighbours | yes
136 |  |  | correction for pseudosymmetry | yes
137 |  |  | GROD angle (grain reference orientation deviation), parent grain reconstruction, ODF (orientation distribution function), separating microstructure into different classes | yes
138 |  |  | dislocation analysis including Burgers vetors and dislocation density | yes
139 |  |  | machine learning |  |  | yes | 100 p | 100 | 100
140 |  | reporting | creating reports from measurement | yes
141 |  |  | template format adjustment |  |  | yes | 10 p | 10 | 10
142 |  | export | export of EBSD patterns as images (e.g. tiff, jpg, bmp) | yes
143 |  | licence | number of off-microscope licenses for the data processing software | 10 | license | yes | 10 p / license | 100 | 10 |  |  | license(-s)
SUPPORT | [9 empty] | 0
elmg field compensation
144 |  | AC and DC compensation |  | yes
servis (including analytics)
145 |  | warranty | warranty of minimum 1 year; optional 2 years | 1 | year | yes | 500 p / 2 year warranty | 500 | 500 |  |  | year(-s)
146 |  | servis | fixed price of service after warranty (years) with inflation increase - price of labor and travel per hour - daily rate (9h + overnight) |  | year | yes | 50 p/ year | 500 | 50 |  |  | year(-s)
147 |  |  | unlimited free software updates | yes
148 |  |  | period of free software upgrades |  | year | yes | 20 p / year | 200 | 20 |  |  | year(-s)
149 |  | delivery | max. period of performance 8 months (předání a uvedení do provozu - handover and comissioning) | 8 | month | yes | 50 p/ month | 100 | 50 |  |  | months
150 |  | training | users traing of minimum 2 days for SEM | 2 | day
151 |  |  | users training of minimum 3 days by application specialist from the producer of EDS-EBSD system (included in the above) | 3 | day