| Confidentiality Level |  | BL - Restricted for internal use | TC ID / Revision |  | 00107408 / D |
| --- | --- | --- | --- | --- | --- |
| Document Status |  | InReviewProcess | Document No. |  | Specific Document code |
| WBS code |  | TBD |  |  |  |
| PBS code |  | TBD |  |  |  |
| Requirements Specification Document (RSD) |  |  |  |  |  |
| Frame contract for Central Vacuum valves TP15_042 Keywords Put keywords related to the document content. |  |  |  |  |  |
|  | Position |  |  | Name |  |
| Responsible person | Group Leader of Vacuum and Cryogenics |  |  | Lukáš Brabec |  |
| Prepared by | Group Leader of Vacuum and Cryogenics |  |  | Lukáš Brabec |  |

| RSS TC ID/revision | Date of Creation | Date of Last Modification | Systems Engineer |
| --- | --- | --- | --- |
| 003842 / A.001 | 10.03.2015 | 10.03.2015 | Aleksei Kuzmenko |
| 003842 / A | 10.03.2015 | 31.03.2015 | Aleksei Kuzmenko |
| 003842 / A | 10.03.2015 | 21.04.2015 | Aleksei Kuzmenko |

| Reviewed By |  |  |  |
| --- | --- | --- | --- |
| Name (Reviewer) | Position | Date | Signature |

| Approved by |  |  |  |
| --- | --- | --- | --- |
| Name (Approver) | Position | Date | Signature |

| Revision History / Change Log |  |  |  |  |
| --- | --- | --- | --- | --- |
| Change No. | Made by | Date | Change description, Pages, Chapters | TC rev. |

Table of Content

1. General description	4

2. International standards	4

3. Technical specification	4

3.1. Product Specific Quality Requirements	5

4. Annex I	6

4.1. Angle valve - SS – pneumatic actuator:	6

4.2. Angle valve - Al – pneumatic actuator:	7

4.3. Angle valve - SS – manual actuator:	8

4.4. Angle valve - Al – manual actuator:	8

4.5. HV Angle valve - SS – pneumatic actuator:	9

4.6. HV Angle valve - Al – pneumatic actuator:	10

4.7. HV Angle valve - SS – manual actuator:	11

4.8. HV Angle valve - Al – manual actuator:	12

4.9. HV Inline valves - SS – pneumatic actuator:	13

4.10. HV Inline valves - Al – pneumatic actuator:	14

4.11. HV Inline valves - SS – manual actuator:	15

4.12. HV Inline valves - Al – manual actuator:	16

4.13. Diaphragm valves – manual actuator:	16

4.14. Gate valve - Al – pneumatic actuator:	17

4.15. Gate valve - SS – pneumatic actuator:	18

4.16. Pendulum valve - Al – pneumatic actuator:	19

4.17. Angle valve with soft-pump function - Al – pneumatic actuator:	20

4.18. UHV angle valve - SS – pneumatic actuator	21

4.19. UHV angle valve - SS – manual actuator	22

4.20. Dosing and Shut-Off valves - manual:	22

4.21. Manually actuated venting valve:	23

4.22. Mini vacuum gate valve for flange KF - Al – pneumatic actuator	23

4.23. Mini UHV gate valve - SS – manual actuator	24

4.24. Gate valve - SS – manual actuator	25

4.25. Pendulum valve - Al – pneumatic actuator – with sapphire window	26

# General description

Vacuum valves will be a part of an extended vacuum system that centrally provides backing and roughing vacuum to large vacuum vessels at ELI-Beamlines facility. The valves will provide sealed separation between two regions with different pressures.

Vacuum valves will be used for applications under high vacuum or ultra-high vacuum.

The vacuum valves are (divided) better grouped or selected by size, type of flange, drive and by the used pressure range.

## Terms, Definitions and Abbreviations

For the purpose of this document, the following abbreviated terms apply:

| Abbreviation | Meaning |
| --- | --- |
| ELI | Extreme Light Infrastructure |
| HV | High Vacuum |
| UHV | Ultra High Vacuum |
| TMP | Turbo Molecular Pump |
| SS | Stainless Steel |
| Al | Aluminium |
| DN | Diameter Nominal |

# International standards

REQ-007697/A

Valves shall be equipped with ISO F flange according to ISO 1609:1986 - Vacuum technology - Flange dimension

REQ-007698/A

Valves shall be equipped with ISO KF flange according to ISO 2861:2013 - Vacuum technology - Dimensions of clamped - type quick-release couplings

REQ-007699/A

Valves shall be equipped with ISO CF flange according to ISO/TS 3669-2:2007 - Vacuum technology – Bakable flanges

# Technical specification

REQ-007700/A

Sizes, types of flanges, actuators and drives shall be according to ANNEX I.

REQ-007701/A

Maximum leak rate shall be according to ANNEX I.

REQ-007702/A

Differential pressure at opening of valves shall be according to ANNEX I.

REQ-007703/A

It shall be possible to install and operate the vacuum valves in any mounting position.

REQ-007704/A

During changing position, the valves shall be particle free.

REQ-007705/A

The electro-pneumatic valves shall be closed (NC-normally closed) during power outage and/or interruption in the supply of compressed air.

REQ-007706/A

The valves shall be equipped with electro-pneumatic actuator with OPEN/CLOSED limit switches and with the possibility of connection to the control system. Connection type shall be specified e.g. DIN 7 pin etc.

REQ-007707/A

Maximum closing or opening time shall be according to ANNEX I.

REQ-007708/A

Valve body material shall be according to ANNEX I.

REQ-007709/A

Valves with pneumatic actuator shall work at (or operate at min) pressure min. 5 bar.

REQ-007710/A

Valves shall be assembled and packed in clean room environment ISO 8 or better according to standard ISO 14644.

REQ-007764/A

Valves sealing surfaces shall be roughness Ra 0.8 μm.

REQ-007903/A

Valves sealing surfaces shall be without scratches.

REQ-008115/A

The sapphire glass shall be in the centre of nominal dimension of valve and according to ANNEX 1, table 4.25.

REQ-008116/A

The material quality of sapphire glass shall be (see ANNEX 1, table 4.25):

- accumulation of micro bubbles (up to 5 bubbles over 50x2 mm);

- 1 microbubble up to 200 microns;

- several bubbles in the edge zone 4 [Kyburz Sapphire].

## Product Specific Quality Requirements

REQ-007814/A

The supplier shall supply Product User Manual.

REQ-007815/A

The supplier shall supply Product Warranty Conditions.

REQ-007816/A

The supplier shall supply CE Declaration of Conformity for each product type (The declaration shall comply with a standards EN ISO/IEC 17050-1; EN ISO/IEC 17050-2).

# Annex I

## Angle valve - SS – pneumatic actuator:

| DN | DN16 | DN25 | DN40 | DN50 | DN63 |
| --- | --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-K (according to ISO 1609) |
| Body shape | Angle | Angle | Angle | Angle | Angle |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 | 0,06 | 0,2 |
| Position indicator | YES | YES | YES | YES | YES |
| Max. closing time [s] | 0,2 | 0,3 | 0,7 | 0,8 | 1 |

Tab. 1 Specification – Angle valves SS – pneumatic actuator

## Angle valve - Al – pneumatic actuator:

| DN | DN16 | DN25 | DN40 | DN50 | DN63 | DN80 | DN100 | DN160 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |  |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) |
| Body shape | Angle | Angle | Angle | Angle | Angle | Angle | Angle | Angle |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 | 0,06 | 0,2 | 0,2 | 0,5 | 1,2 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time [s] | 0,2 | 0,3 | 0,7 | 0,8 | 1 | 1 | 1,5 | 3 |

Tab. 2 Specification – Angle valves Al – pneumatic actuator

## Angle valve - SS – manual actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Angle | Angle | Angle | Angle |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 | 8000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |  |

Tab. 3 Specification – Angle valves SS – manual actuator

## Angle valve - Al – manual actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Angle | Angle | Angle | Angle |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 | 8000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |  |

Tab. 4 Specification – Angle valves Al – manual actuator

## HV Angle valve - SS – pneumatic actuator:

| DN | DN16 | DN25 | DN40 | DN50 | DN63 | DN100 | DN160 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) |
| Body shape | Angle | Angle | Angle | Angle | Angle | Angle | Angle |
| Feedthrough | Bellows | Bellows | Bellows | Bellows | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 | 0,06 | 0,2 | 0,5 | 1,2 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time [s] | 0,2 | 0,3 | 0,7 | 0,8 | 1 | 1,5 | 3 |
| Max. particles generation of the valve [nm] | 20 |  |  |  |  |  |  |

Tab. 5 Specification – HV Angle valves SS – pneumatic actuator

## HV Angle valve - Al – pneumatic actuator:

| DN | DN16 | DN25 | DN40 | DN50 | DN63 | DN80 | DN100 | DN160 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |  |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) |
| Body shape | Angle | Angle | Angle | Angle | Angle | Angle | Angle | Angle |
| Feedthrough | Bellows | Bellows | Bellows | Bellows | Bellows | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 | 0,06 | 0,2 | 0,2 | 0,5 | 1,2 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time [s] | 0,2 | 0,3 | 0,7 | 0,8 | 1 | 1 | 1,5 | 3 |
| Max. particles generation of the valve [nm] | 20 |  |  |  |  |  |  |  |

Tab. 6 Specification – HV Angle valves Al – pneumatic actuator

## HV Angle valve - SS – manual actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Angle | Angle | Angle | Angle |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 | 8000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |  |
| Max. particles generation of the valve [nm] | 20 |  |  |  |

Tab. 7 Specification – HV Angle valves SS – manual actuator

## HV Angle valve - Al – manual actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Angle | Angle | Angle | Angle |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 | 8000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |  |
| Max. particles generation of the valve [nm] | 20 |  |  |  |

Tab. 8 Specification – HV Angle valves Al – manual actuator

## HV Inline valves - SS – pneumatic actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Inline | Inline | Inline | Inline |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 | 0,06 |
| Position indicator | YES | YES | YES | YES |
| Max. closing time[s] | 0,2 | 0,3 | 0,7 | 0,8 |
| Max. particles generation of the valve [nm] | 20 |  |  |  |

Tab. 9 Specification – HV Inline valves SS – pneumatic actuator

## HV Inline valves - Al – pneumatic actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Inline | Inline | Inline | Inline |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 | 0,06 |
| Position indicator | YES | YES | YES | YES |
| Max. closing time[s] | 0,2 | 0,3 | 0,7 | 0,8 |
| Max. particles generation of the valve [nm] | 20 |  |  |  |

Tab. 10 Specification – HV Inline valves Al – pneumatic actuator

## HV Inline valves - SS – manual actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Inline | Inline | Inline | Inline |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 | 8000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |  |
| Max. particles generation of the valve [nm] | 20 |  |  |  |

Tab. 11 Specification – HV Inline valves SS – manual actuator

## HV Inline valves - Al – manual actuator:

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Inline | Inline | Inline | Inline |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 | 8000 |
| Max leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |  |
| Max. particles generation of the valve [nm] | 20 |  |  |  |

Tab. 12 Specification – HV Inline valves Al – manual actuator

## Diaphragm valves – manual actuator:

| DN | DN16 | DN25 | DN40 |
| --- | --- | --- | --- |
| Body material | Aluminium |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Feedthrough | Diaphragm | Diaphragm | Diaphragm |
| Minimal cycles until first service | 15000 | 15000 | 15000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 |

Tab. 13 Specification – Diaphragm valves – manual actuator

## Gate valve - Al – pneumatic actuator:

| DN | DN63 | DN80 | DN100 | DN160 | DN200 | DN250 | DN320 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |  |  |  |
| Flange | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) |
| Body shape | Gate | Gate | Gate | Gate | Gate | Gate | Gate |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Minimal cycles until first service | 150000 | 150000 | 150000 | 80000 | 80000 | 80000 | 80000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,3 | 0,4 | 0,4 | 2 | 3 | 3 | 3,5 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time[s] | 2,5 | 2,5 | 3 | 4 | 8 | 8 | 10 |
| Max. particles generation of the valve [nm] | 150 |  |  |  |  |  |  |

Tab. 14 Specification – Gate valves Al – pneumatic actuator

These above 4.14 specified gate valves shall be compatible with TMP mounting.

## Gate valve - SS – pneumatic actuator:

| DN | DN63 | DN80 | DN100 | DN160 | DN200 | DN250 | DN320 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |  |  |  |
| Flange | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) |
| Body shape | Gate | Gate | Gate | Gate | Gate | Gate | Gate |
| Feedthrough | Bellows | Bellows | Bellows | Bellows | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Minimal cycles until first service | 150000 | 150000 | 150000 | 150000 | 150000 | 150000 | 150000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,2 | 0,2 | 0,3 | 0,3 | 0,5 | 0,5 | 0,7 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time[s] | 2,5 | 2,5 | 3 | 4 | 8 | 8 | 10 |
| Max. particles generation of the valve [nm] | 1200 |  |  |  |  |  |  |

Tab. 15 Specification – Gate valves SS – pneumatic actuator

These valves must be possible used for TMP mounting.

## Pendulum valve - Al – pneumatic actuator:

| DN | DN100 | DN160 | DN200 | DN250 | DN320 | DN400 | DN500 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |  |  |  |
| Flange | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) |
| Body shape | Pendulum | Pendulum | Pendulum | Pendulum | Pendulum | Pendulum | Pendulum |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 25 | 8 | 2 | 2 | 2 | 2 | 2 |
| Minimal cycles until first service | 150000 | 150000 | 150000 | 150000 | 150000 | 150000 | 80000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,4 | 0,4 | 0,5 | 0,5 | 0,8 | 1 | 4 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time[s] | 5 | 5 | 6 | 6 | 8 | 8 | 20 |
| Max. particles generation of the valve [nm] | 1200 |  |  |  |  |  |  |

Tab. 16 Specification – Pendulum valves Al – pneumatic actuator

These valves shall be possible used for TMP mounting.

Valve DN500 shall be possible to use for Cryopump mounting.

## Angle valve with soft-pump function - Al – pneumatic actuator:

| DN | DN25 | DN40 | DN50 | DN63 | DN80 | DN100 | DN160 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) |
| Body shape | Angle | Angle | Angle | Angle | Angle | Angle | Angle |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,03 | 0,05 | 0,07 | 0,25 | 0,25 | 0,5 | 1,2 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time[s] | 0,35 | 0,8 | 0,9 | 1 | 1 | 1,5 | 3,5 |

Tab. 17 Specification - Soft start valves Al – pneumatic actuator

Solenoids shall be 24V DC

## UHV angle valve - SS – pneumatic actuator

| DN | DN16 | DN40 | DN63 |
| --- | --- | --- | --- |
| Body material | Stainless steel |  |  |
| Flange | CF-R (according to ISO/TS 3669) | CF-R (according to ISO/TS 3669) | CF-R (according to ISO/TS 3669) |
|  | Rotatable CF flange | Rotatable CF flange | Rotatable CF flange |
| Body shape | Angle | Angle | Angle |
| Solenoid valve [VDC] | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 2∙106 | 2∙106 | 2∙106 |
| Max leak rate – valve seat [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,05 | 0,2 |
| Position indicator | YES | YES | YES |
| Max. closing time[s] | 0,2 | 0,7 | 1 |
| Max. particles generation of the valve [nm] | 20 |  |  |

Tab. 18 Specification – UHV angle valves SS – pneumatic actuator

## UHV angle valve - SS – manual actuator

| DN | DN16 | DN40 | DN63 |
| --- | --- | --- | --- |
| Body material | Stainless steel |  |  |
| Flange | CF-R (according to ISO/TS 3669) | CF-R (according to ISO/TS 3669) | CF-R (according to ISO/TS 3669) |
|  | Rotatable CF flange | Rotatable CF flange | Rotatable CF flange |
| Body shape | Angle | Angle | Angle |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 10000 | 10000 | 10000 |
| Max leak rate – valve seat [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |
| Max. particles generation of the valve [nm] | 20 |  |  |

Tab. 19 Specification – UHV angle valves SS – manual actuator

## Dosing and Shut-Off valves - manual:

| DN | DN16 |
| --- | --- |
| Body material | Stainless steel |
| Flange | ISO-KF (according to ISO 2861) |
| Body shape | Angle |
| Min. pressure range [bar] | 1 (abs) |
| Max leak rate [mbar ls-1] | 5∙10-9 |

Tab. 20 Specification – Dosing and Shut-Off valves

## Manually actuated venting valve:

| DN | DN10 |
| --- | --- |
| Body material | Stainless steel |
| Flange | ISO-KF (according to ISO 2861) |
| Min. differential pressure on plate in opening direction [mbar] | 1000 |
| Max leak rate – valve seat [mbar ls-1] | 5∙10-9 |

Tab. 21 Specification – Manually actuated venting valves

## Mini vacuum gate valve for flange KF - Al – pneumatic actuator

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Gate | Gate | Gate | Gate |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 25 | 25 | 25 | 25 |
| Minimal cycles until first service | 30000 | 30000 | 30000 | 30000 |
| Max. leak rate [mbar ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,05 | 0,08 | 0,1 | 0,1 |
| Position indicator | YES | YES | YES | YES |
| Max. closing time [s] | 1 | 1,5 | 1,8 | 2 |

Tab. 22 Specification – Mini vacuum gate valve for KF Al – pneumatic actuator

## Mini UHV gate valve - SS – manual actuator

| DN | DN40 |
| --- | --- |
| Body material | Stainless steel |
| Flange | CF-F (according to ISO/TS 3669) Metric threads |
| Body shape | Gate |
| Feedthrough | Bellows |
| Min. differential pressure at opening [mbar] | 25 |
| Minimal cycles until first service | 20000 |
| Max leak rate – valve seat [mbar ls-1] | 5∙10-9 |
| Position indicator | YES |

Tab. 23 Specification – Mini UHV gate valve SS – manual actuator

## Gate valve - SS – manual actuator

| DN | DN63 |
| --- | --- |
| Body material | Stainless steel |
| Flange | CF-F (according to ISO/TS 3669) Metric threads |
| Body shape | Gate |
| Feedthrough | Bellows |
| Min. differential pressure at opening [mbar] | 25 |
| Minimal cycles until first service | 100000 |
| Max leak rate – valve seat [mbar ls-1] | 5∙10-9 |
| Position indicator | YES |
| Max. particles generation of the valve [nm] | 1200 |

Tab. 24 Specification – Gate valve SS – manual actuator

## Pendulum valve - Al – pneumatic actuator – with sapphire window

| DN | DN400 |
| --- | --- |
| Body material | Aluminium |
| Flange | ISO-F (according to ISO 1609) |
| Body shape | Pendulum |
| Solenoid valve [VDC] | 24 |
| Min. differential pressure at opening [mbar] | 2 |
| Minimal cycles until first service | 150000 |
| Max. leak rate [mbar∙l∙s-1] | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 1 |
| Position indicator | YES |
| Max. closing time [s] | 8 |
| Max. particles generation of the valve [nm] | 1200 |
| Sapphire window |  |
| Type of window | sapphire glass (see REQ-008116/A) |
| Optical φ [mm] | 280 |
| Parallelism max. | 3’ |
| Surface finish | both sides optically polished 60/40 S/D |
| Orientation | Random |
| Wavelength range | Visible |
| Coatings | NON |

Tab. 25 Specification – Pendulum valves Al – pneumatic actuator – with sapphire window