# Annex no. 1 – Technical specification

Public Contract Name:

“High power ps fiber laser for material processing”

Technical specification of galvanometric scanning system

|  | The subject of the procurement | High-power pulsed picosecond fiber laser for material processing |
| --- | --- | --- |
|  | Following tasks to be realized with galvanometric scanning system | - High-precision laser micromachining and material surface structuring. |
|  |  | - Ready for 24/7 operation |
|  |  | - Compatible with multi-beam and beamshaping technologies based on spatial light modulator |
|  |  | - Fundamental wavelength in 1 µm region |
|  |  | - Adjustable pulse repetition rate and pulse duration to optimize for different materials and processes. |
|  |  | - Compatible with Direct Machining Control system |
|  |  | - Integration with existing galvanometric scanning systems and RTC control boards. |
| Technical features of galvanometric scanning system |  |  |
| Optical Characteristics |  | - Wavelength: Must be between 1020 nm and 1070 nm. - Pulse Repetition Rate: Must be adjustable at least in the range of 100 kHz and 5000 kHz - Pulse Duration: Must be tunable within the range of at least 1 to 3 picoseconds. - Output Power: Maximum output power should be at least 90 W - Beam Quality (M²): Must be lower than 1.3. - Polarization: Must be linear, with a polarization extinction ratio of at least 20 dB. |
| Performance and Stability |  | - Pulse Energy: Maximum pulse energy must be at least 100 µJ at lower repetition rates. - Power Stability: Output power stability should be within 1% over 8 hours of continuous operation. - Pulse-to-Pulse Stability: RMS < 2% at the nominal pulse repetition rate. |
| Control and Interfaces |  | - Must be compatible with both analog and TTL control signals, with support for RS-232 or Ethernet communication. - The system must include an integrated pulse picker for flexible pulse shaping. - Control software must be compatible with Windows 11 and support integration with existing systems (e.g., Scanlab’s laserDesk and RTC control, Direct Machining Control). |
| Dimensions and Weight |  | Control unit dimensions should not exceed 500 x 650 x 150 mm, with a weight limit of 40 kg. |
| Accessories |  | Delivery with all necessary cables, connectors, and manuals in English. Must come with power supply for 24/7 operation. Must come with dedicated control software for laser operation |