Technical specification of GHz burst repetition rate femtosecond laser

|  | The subject of the procurement | GHz burst repetition rate femtosecond laser |
| --- | --- | --- |
|  | Following tasks to be realized with galvanometric scanning system | - High-precision laser micromachining and volumetric modifications in transparent materials. |
|  |  | - Ready for 24/7 operation |
|  |  | - The system should generate 1030 nm output beam with possibilities of second (2H), third (3H), and fourth (4H) harmonic generations |
|  |  | - Adjustable burst mode up to GHz region for applications such as selective ablation, deep engraving, surface polishing, and functionalization |
|  |  | - Adjustable pulse repetition rate and pulse duration to optimize for different materials and processes. |
|  |  | - Compatible with Direct Machining Control system |
|  |  | - Integration with existing galvanometric scanning systems and RTC control boards. |
| Technical features of GHz burst repetition rate femtosecond laser |  |  |
| Optical Characteristics |  | - Wavelengths: - Fundamental: 1030 nm - Second harmonic (2H): 515 nm - Third harmonic (3H): 343 nm - Fourth harmonic (4H): 257 nm - Pulse Repetition Rate: must be adjustable up to 2 MHz, with burst mode options up to GHz region - Pulse Duration: Must be tunable tunable at least in the range of 250 fs and 1 ps. - Output Power: Maximum output power of at least 120W at 1030 nm with > 50% efficiency for the second harmonic and > 20% for the third harmonic. - Burst Mode Capability: Must support burst mode with both nanosecond and picosecond pulse separation, configurable to create a set of pulse sub-packets for enhanced control in material processing. - Beam Quality (M²): Must be lower than 1.2. - Polarization: Must be vertical for the fundamental wavelength |
| Performance and Stability |  | - Power Stability: Output power stability should be within 1% over 8 hours of continuous operation. - Pulse Energy Stability: Stability of pulse energy must be <0.5% (RMS). |
| Control and Interfaces |  | - Must be compatible with analog, TTL, and Ethernet-based controls. - The system should provide interfaces for external synchronization and control, supporting flexible repetition rate tuning and burst shaping - Must be supplied with control software that allows adjustment of pulse repetition rates, burst configurations, and harmonic switching. |
| Accessories |  | - Delivery with all necessary cables, connectors, and manuals in English. - The system must include all necessary optical components for harmonic generation (2H, 3H, 4H) and mounting fixtures. Include power supply, cabling, and other connectors required for operation. |