| Confidentiality Level |  | BL - Restricted for internal use | TC ID / Revision | 00167797/C |
| --- | --- | --- | --- | --- |
| Document Status |  | Document Released | Document No. | N/A |
| WBS code |  | 6.3 – Vacuum Engineering |  |  |
| PBS code |  | NOT known at the moment. The vacuum valves for L4, Beam Transport and Experimental halls will be delivered within this framework contract. The future PBS codes will include following top level strings: RA.L4, SE.BDS and E.E[x] |  |  |
| Project branch |  | Engineering & Scientific documents (E&S) |  |  |
| Document Type |  | Specification (SP) |  |  |
| [RSD product category A] Framework contract for Vacuum valves in 2018 (TP18_034) Keywords vacuum, valves, particles, flange |  |  |  |  |
|  | Position |  |  | Name |
| Responsible person | Group Leader of Vacuum and Cryogenics |  |  | Lukáš Brabec |
| Prepared by | Group Leader of Vacuum and Cryogenics |  |  | Lukáš Brabec |

| RSS TC ID/revision | RSS - Date of Creation | RSS - Date of Last Modification | Systems Engineer |
| --- | --- | --- | --- |
| 013904/A.001 | 24.04.2018 16:51 | 03.05.2018 17:18 | David Myslikovjan |
| 013904/A.002 | 04.05.2018 16:39 | 04.05.2018 17:37 | David Myslikovjan |
| 013904/A.003 | 16.05.2018 15:45 | 16.05.2018 15:45 | David Myslikovjan |

| Approved by |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Name (Approver) |  |  | Position |  | Date |  | Signature |  |  |  |
| Martin Laub |  |  | Chief Engineer |  |  |  |  |  |  |  |
| Reviewed By |  |  |  |  |  |  |  |  |  |
| Name (Reviewer) |  | Position |  |  |  | Date |  | Signature |  |
| Luboš Nims |  | Head of Electrical Engineering |  |  |  | NOTICE |  |  |  |
| Jiří Trdlička |  | Control System Engineer |  |  |  | NOTICE |  |  |  |
| Ladislav Půst |  | Manager installation of technology |  |  |  | NOTICE |  |  |  |
| Petr Vobořil |  | Designer |  |  |  | NOTICE |  |  |  |
| Veronika Olšovcová |  | Safety Coordinator |  |  |  | NOTICE |  |  |  |
| Viktor Fedosov |  | Quality Manager |  |  |  | NOTICE |  |  |  |
| Martin Laub |  | Chief Engineer |  |  |  | NOTICE |  |  |  |
| Revision History / Change Log |  |  |  |  |  |  |  |  |  |  |
| Change No. | Made by |  | Date | Change description, Pages, Chapters |  |  |  |  | TC rev. |  |
| 1 | L. Brabec, D. Myslikovjan |  | 30.04.2018 | RSD draft creation |  |  |  |  | A |  |
| 2 | D. Myslikovjan |  | 04.05.2018 | RSD update, version for internal review |  |  |  |  | B |  |
| 3 | D. Myslikovjan |  | 16.05.2018 | RSD approved, final version |  |  |  |  | C |  |

Table of Content

1. Introduction	4

1.1. Purpose	4

1.2. General description	4

1.3. Scope	4

1.4. Terms, Definitions and Abbreviations	5

1.5. Reference documents	5

2. General Functional, Performance and Design requirements	5

3. Transportation requirements	6

4. Product Safety Requirements	7

5. Product Quality Requirements	7

5.1. International standards	8

6. ANNEX I	9

6.1. HV Angle valve - Al – pneumatic actuator	9

6.2. HV Angle valve - Al – manual actuator	10

6.3. HV Inline valves - Al – pneumatic actuator	11

6.4. HV Inline valves - Al – manual actuator	12

6.5. Gate valve - Al – pneumatic actuator	13

6.6. Gate valve - SS – pneumatic actuator	14

6.7. Angle valve with soft-pump function - Al – pneumatic actuator	15

6.8. Manually actuated venting valve	16

6.9. Mini vacuum gate valve for flange KF - Al – pneumatic actuator	17

# Introduction

## Purpose

This Requirements Specification Document (RSD) lists the technical requirements and constraints on product/device (Vacuum valves) of ELI-Beamlines project. This can lead to the identification of product interfaces with the ELI science based technology and ELI building facility. This RSD also acts as the parent document for the technical requirements that need to be addressed in lower level design description documents.

## General descri ption

Vacuum valves will be a part of an extended vacuum system that centrally provides backing and roughing vacuum to large vacuum vessels at ELI-Beamlines facility. The valves will provide sealed separation between two regions with different pressures.

Vacuum valves will be used for applications under high vacuum or ultra-high vacuum.

The vacuum valves categorized according to type of flange, drive and applicable pressure range.

## Scope

This RSD contains all of the technical requirements: functional, performance, operational and design, transportation, safety and quality requirements for the product/device Vacuum valves. This RSD and contract are a follow-up to the previous valves-related framework contracts which are almost expired. The PBS codes are not known at the moment. Nevertheless it is known that the vacuum valves for laser L4, Beam Transport and Experimental halls will be delivered within this framework contract. The future PBS codes will include following top level strings: RA.L4, SE.BDS and E.E[x] where the "x" parameter can be substituted by the respective number of ELI experimental hall(s).

This product is a product Category A according to the ELI Beamlines RSD categories. The category A is an Off-the-shelf Product without necessity of modifications and necessity to be subjected to a verification programme (review of design, inspection and testing) for ELI applications by the actual project specifications. All verification activities performing by a Supplier shall be executed in accordance with the Supplier’s plan of outgoing inspection and tests. Internal Acceptance Procedure of the product Category A shall be established and applied before the product implementation (operation phase).

## Terms, Definitions and Abbreviations

For the purpose of this document, the following abbreviated terms apply:

| Abbreviation | Meaning |
| --- | --- |
| ELI | Extreme Light Infrastructure |
| HV | High Vacuum |
| UHV | Ultra High Vacuum |
| TMP | Turbo Molecular Pump |
| SS | Stainless Steel |
| Al | Aluminium |
| DN | Diameter Nominal |
| CA | Contract Authority |

## Reference documents

| Doc number | Title of document |
| --- | --- |
| RD-01 | 00163567-B_1.2_Q_M_Guide_for_Instructions_for_Use_Ver-9_EN_fully_signed |

# General Functional, Performance and Design requirements

REQ-023170/A

Sizes, types of flanges, actuators and drives shall be according to ANNEX I.

REQ-023171/A

Maximum leak rate shall be according to ANNEX I.

REQ-023172/A

Differential pressure at opening of valves shall be according to ANNEX I.

REQ-023173/A

It shall be possible to install and operate the vacuum valves in any mounting position.

REQ-023174/A

Gate valves specified in chapter 6.5 of ANNEX I (Gate valve - Al – pneumatic actuator) shall be compatible with TMP mounting.

REQ-023175/A

Gate valves specified in chapter 6.6 of ANNEX I (Gate valve - SS – pneumatic actuator) shall be compatible with TMP mounting.

REQ-023176/A

All solenoids valves specified in ANNEX I shall use 24V DC power supply.

REQ-023177/A

The maximal amount of particles (particles bigger than 300 nm) generation of the valve during one cycle shall be according to 
ANNEX I.

REQ-023178/A

The electro-pneumatic valves shall be closed (NC-normally closed) during power outage and/or interruption in the supply of compressed air.

REQ-023179/A

The valves shall be equipped with electro-pneumatic actuator with OPEN/CLOSED limit switches and with the possibility of connection to the control system. Connection type shall be specified e.g. DIN 7 pin etc.

REQ-023180/A

Maximum closing or opening time shall be according to ANNEX I.

REQ-023181/A

Valve body material shall be according to ANNEX I.

REQ-023182/A

Valves with pneumatic actuator shall work at (or operate at min) pressure min. 5 bar.

REQ-023183/A

Valves shall be assembled and packed in clean room environment ISO 8 according to standard ČSN EN ISO 14644.

REQ-023184/A

Valves sealing surfaces shall be roughness Ra 0.8 μm.

REQ-023185/A

Valves sealing surfaces shall be without scratches.

# Transportation requirements

REQ-023186/A

The transportation to the final destination of the technologies and the instruments shall be conducted by the supplier.

REQ-023187/A

The technologies and instruments shall be delivered in protective package preventing damage and contamination and a minimum of two plies separate clean packaging. The technologies shall be cleaned and packaged, in compliance with the cleanliness of class 8 according to ČSN EN ISO 14644.

# Product Safety Requirements

REQ-023188/A

The Supplier shall supply a Declaration of Conformity or any other

equivalent document legally recognized and accepted in the Czech

Republic for each product type if the appropriate legislation

determines the Supplier’s obligation to have a Declaration of

Conformity (or the equivalent document) for the purposes of a

Product sale in the Czech Republic to fulfil the requirements of

2001/95/EC directive or applicable Czech law.

# Product Quality Requirements

REQ-023189/A

The Supplier shall provide Instructions for use (Product User Manual) as part of the delivered Product. The Instructions for use shall be approved by the CA and shall include the instructions and descriptions regarding the following:

- transport;

- handling;

- storage;

- installation;

- safe operation and maintenance procedures.

The Instructions for use shall be written in accord with ČSN EN 82079-1 (equivalent to IEC - 82079-1:2012).

Note: As an alternative to standard ČSN EN 82079-1 (equivalent to IEC - 82079-1:2012) can be used internal ELI guide - see RD-01 document reference.

REQ-023190/A

The Supplier shall provide with the delivery of the Product (aside from Instructions for use) the following types of documents:

•	the manufacturing drawings (3D model and 2D drawings) in electronic form as defined in REQ-023195/A.

REQ-023195/A

The Supplier shall use following data formats:

•	*.JPG, *.PNG, *.PDF/A, *.HTML;

•	CAD 2D: *.dwg;

•	CAD 3D: STEP file types (*.stp;*.ste;*.step);

•	text processors *.doc, *.docx, OpenDocument Format;

•	spreadsheet processors *.xls, *.xlsx, OpenDocument Format;

•	presentations *.ppt, *.pptx; OpenDocument Format.

REQ-023191/A

The Supplier shall provide information on executed of outgoing check of the Product/Device. At least this information shall comprise declaration about execution of outgoing check and declaration of conformity with technical requirements defined by the product RSD and completeness of the Product/Device.

NOTE: Alternatively the Supplier might provide the CA with information (e.g.: catalogue/technical data sheets, product manuals or other similar documentation) subject to the condition that such documentation is detailed enough to prove meeting all requirements stipulated herein.

REQ-023192/A

The Supplier shall establish and maintain a non-conformance control system compatible with ČSN EN ISO 9001 (equivalent to EN ISO 9001).

## International standards

REQ-023193/A

Valves shall be equipped with ISO F flange according to ISO 1609:1986 - Vacuum technology - Flange dimension.

REQ-023194/A

Valves shall be equipped with ISO KF flange according to ISO 2861:2013 - Vacuum technology - Dimensions of clamped - type quick-release couplings.

# ANNEX I

## HV Angle valve - Al – pneumatic actuator

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Angle | Angle | Angle | Angle |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar. ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 | 0,06 |
| Position indicator | YES | YES | YES | YES |
| Max. closing time [s] | 0,2 | 0,3 | 0,7 | 0,8 |
| Max. particles generation of the valve | 20 |  |  |  |

Tab. 1 Specification – HV Angle valves Al – pneumatic actuator

## HV Angle valve - Al – manual actuator

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Angle | Angle | Angle | Angle |
| Feedthrough | Bellows | Bellows | Bellows | Bellows |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 | 8000 |
| Max. leak rate [mbar. ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |  |
| Max. particles generation of the valve | 20 |  |  |  |

Tab. 2: Specification – HV Angle valves Al – manual actuator

## HV Inline valves - Al – pneumatic actuator

| DN | DN16 | DN25 | DN40 |
| --- | --- | --- | --- |
| Body material | Aluminium |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Inline | Inline | Inline |
| Feedthrough | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar. ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,02 | 0,025 | 0,05 |
| Position indicator | YES | YES | YES |
| Max. closing time[s] | 0,2 | 0,3 | 0,7 |
| Max. particles generation of the valve | 20 |  |  |

Tab. 3: Specification – HV Inline valves Al – pneumatic actuator

## HV Inline valves - Al – manual actuator

| DN | DN16 | DN25 | DN40 |
| --- | --- | --- | --- |
| Body material | Aluminium |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Inline | Inline | Inline |
| Feedthrough | Bellows | Bellows | Bellows |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 8000 | 8000 | 8000 |
| Max leak rate [mbar. ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Handwheel | Removable |  |  |
| Max. particles generation of the valve | 20 |  |  |

Tab. 4: Specification – HV Inline valves Al – manual actuator

## Gate valve - Al – pneumatic actuator

| DN | DN63 | DN80 | DN100 | DN160 | DN200 | DN250 |
| --- | --- | --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |  |  |
| Flange | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) |
| Body shape | Gate | Gate | Gate | Gate | Gate | Gate |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 25 | 25 | 25 | 25 | 25 | 25 |
| Minimal cycles until first service | 150000 | 150000 | 150000 | 80000 | 80000 | 80000 |
| Max. leak rate [mbar. ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,3 | 0,4 | 0,4 | 2 | 3 | 3 |
| Position indicator | YES | YES | YES | YES | YES | YES |
| Max. closing time[s] | 2,5 | 2,5 | 3 | 4 | 8 | 8 |
| Max. particles generation of the valve | 150 |  |  |  |  |  |

Tab. 5: Specification – Gate valves Al – pneumatic actuator

## Gate valve - SS – pneumatic actuator

| DN | DN63 | DN80 | DN100 | DN160 | DN200 | DN250 | DN320 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Stainless steel |  |  |  |  |  |  |
| Flange | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) | ISO-F (according to ISO 1609) |
| Body shape | Gate | Gate | Gate | Gate | Gate | Gate | Gate |
| Feedthrough | Bellows / rotary | Bellows / rotary | Bellows | Bellows | Bellows | Bellows | Bellows |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Minimal cycles until first service | 150000 | 150000 | 150000 | 150000 | 150000 | 150000 | 150000 |
| Max. leak rate [mbar.ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,2 | 0,2 | 0,3 | 0,3 | 0,5 | 0,5 | 0,7 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time[s] | 2,5 | 2,5 | 3 | 4 | 8 | 8 | 10 |
| Max. particles generation of the valve | 1200 |  |  |  |  |  |  |

Tab. 6: Specification – Gate valves SS – pneumatic actuator

## Angle valve with soft-pump function - Al – pneumatic actuator

| DN | DN25 | DN40 | DN50 | DN63 | DN80 | DN100 | DN160 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) | ISO-K (according to ISO 1609) |
| Body shape | Angle | Angle | Angle | Angle | Angle | Angle | Angle |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 |
| Minimal cycles until first service | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 | 9∙105 |
| Max. leak rate [mbar. ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,03 | 0,05 | 0,07 | 0,25 | 0,25 | 0,5 | 1,2 |
| Position indicator | YES | YES | YES | YES | YES | YES | YES |
| Max. closing time[s] | 0,35 | 0,8 | 0,9 | 1 | 1 | 1,5 | 3,5 |
| Max. particles generation of the valve | 150 |  |  |  |  |  |  |

Tab. 7: Specification - Soft start valves Al – pneumatic actuator

## Manually actuated venting valve

| DN | DN10 |
| --- | --- |
| Body material | Stainless steel |
| Flange | ISO-KF (according to ISO 2861) |
| Min. differential pressure on plate in opening direction [mbar] | 1000 |
| Max leak rate – valve seat [mbar. ls-1] | 5∙10-9 |

Tab. 8: Specification – Manually actuated venting valves

## Mini vacuum gate valve for flange KF - Al – pneumatic actuator

| DN | DN16 | DN25 | DN40 | DN50 |
| --- | --- | --- | --- | --- |
| Body material | Aluminium |  |  |  |
| Flange | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) | ISO-KF (according to ISO 2861) |
| Body shape | Gate | Gate | Gate | Gate |
| Feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough | Shaft feedthrough |
| Solenoid valve [VDC] | 24 | 24 | 24 | 24 |
| Min. differential pressure at opening [mbar] | 25 | 25 | 25 | 25 |
| Minimal cycles until first service | 30000 | 30000 | 30000 | 30000 |
| Max. leak rate [mbar. ls-1] | 5∙10-9 | 5∙10-9 | 5∙10-9 | 5∙10-9 |
| Max. consumption of compressed air [l/stroke] | 0,05 | 0,08 | 0,1 | 0,1 |
| Position indicator | YES | YES | YES | YES |
| Max. closing time [s] | 1 | 1,5 | 1,8 | 2 |
| Max. particles generation of the valve | 150 |  |  |  |

Tab. 9: Specification – Mini vacuum gate valve for KF Al – pneumatic actuator